## data sheet
LGR 10-POINT CALIBRATION CHECK DATA SHEET
SESSION INFO | [6 empty] | STANDARD INFORMATION
LGR Serial Number: |  |  | US4300160800001173 |  |  |  |  | CH4 (ppm) | CO2 (ppm) | Unique identifier (i.e. lot number)
Date: | 2019-04-26 | [5 empty] | low | 1.75 | 302 | Praxair Batch 70001719404
Analyst: | AB + DD
Laser A τ: | [6 empty] | high | 504.3 | 506 | Liquide reference 54-401447400-1
Laser B τ:
File Name:
STANDARD CALCULATIONS | [5 empty] | MEASURED VALUES
Point | Slow | Shigh | total flow | CH4 known | CO2 known | Start Time: | End Time: | CH4 | CO2
1 | 0.9 | 0 | 0.9 | 1.75 | 302 | 15:31 | 15:36 | 1.765 | 295.65
2 | 4 | 0.05 | 4.05 | 7.9543 | 304.5185 | 15:39 | 15:43 | 7.58 | 297.09
3 | 3 | 0.05 | 3.05 | 9.9885 | 305.3443 | 15:43 | 15:47 | 9.4903 | 298.79
4 | 2 | 0.05 | 2.05 | 14.0073 | 306.9756 | 15:47 | 15:51 | 13.6814 | 299.63
5 | 2 | 0.1 | 2.1 | 25.681 | 311.7143 | 15:51 | 15:57 | 24.6137 | 304.93
6 | 1.5 | 0.1 | 1.6 | 33.1594 | 314.75 | 15:57 | 16:00 | 32.0708 | 307.36
7 | 1.5 | 0.2 | 1.7 | 60.8735 | 326 | 16:00 | 16:04 | 56.258 | 318.41
8 | 1 | 0.25 | 1.25 | 102.26 | 342.8 | 16:04 | 16:08 | 92.58 | 337.301
9 | 0.7 | 1 | 1.7 | 297.3676 | 422 | 16:10 | 16:13 | 261.0754 | 414.12
10 | 0 | 0.9 | 0.9 | 504.3 | 506 | 16:14 | 16:18 | 439.4305 | 496.95

## dropdown list
15-0016
US4300160800001173